## Slide 1: Routine Analysis of Soils using ICP-MS and Discrete Sampling
Agilent 7850 ICP-MS provides the stability and robustness needed for the long-term accurate analysis of high matrix samples
DE44334.7416666667

## Slide 2: Routine Analysis of Soils – complex and high matrix samples
The Challenge
Routine monitoring of environmental samples is required to comply with regulations in many countries. But samples such as soils, sediments, and sludges can be problematic:
Complex and high matrices (high % of total dissolved solids)
Concentrations of some elements are high (often % level)
Analytes must be measured accurately from low (trace elements) to high concentration (major elements)
The Solution
Analysis requires ICP-MS with good matrix tolerance, effective control of matrix-based interferences, and wide dynamic range.
High matrix sample measurements have a potential for signal instability due to matrix deposition, and frequent maintenance requirements.

## Slide 3: Environmental Testing Laboratories Face…
Preparing calibration standards and samples
Developing methods
Performing daily checks
Reviewing data
Reporting results
So, labs can benefit from reducing the time required for key steps in the workflow:
Challenge of managing many samples
Development of new methods to meet evolving regulations
Ensuring data quality
Maintaining a high level of productivity and data quality is vital to the success of many labs.

## Slide 4: Benefits of the Agilent 7850 ICP-MS hardware
Easy to setup and use
    Method Wizard, Preset Methods and Autotune, M2+ correction
Effective ORS4 helium collision mode
    Effective removal of common polyatomic ion interferences
10 orders dynamic range
    Enables the measurement of major and trace elements in the same run
Robust, matrix tolerant plasma (low CeO+/Ce+)
    Gives consistent results in a wide range of different sample types
High sample throughput
    Shortens time for sample introduction and rinse
Minimizing carry over from previous samples
    Sample is delivered to the plasma only during the actual data acquisition step
Benefits of the ISIS 3 discrete sampling system

## Slide 5: Benefits of ICP-MS MassHunter Software
Easy to setup and use
    Method Wizard, Preset Methods and Autotune, M2+ correction
IntelliQuant
    Total Matrix Solid (TMS) info, Heat map display of data in periodic table view
Automatic maintenance alerts
    Early Maintenance Feedbacks (EMF), Post-run check
This unique combination of hardware performance and software functionality is one of the reasons that hundreds of laboratories worldwide use Agilent ICP-MS.
EMF
Post-run TuneCheck
Heat map

## Slide 6: Experimental
Instrument
  Agilent 7850 ICP-MS
  ISIS 3 discrete sampling system
  SPS 4 autosampler
Plasma mode
  UHMI-8
Lens voltage
  Autotune
Cell mode
  Helium *
  *Enhanced He mode settings used for Se and P. As and Se were measured at the increased resolution settings used for M2+ Correction
ICP-MS Parameter | He mode | 
Plasma mode | UHMI-8 | 
RF power (W) | 1600 | 
Spray chamber temp (°Ｃ) | 2 | 
Sampling Depth (mm) | 10 | 
Nebulizer gas flow (L/min) | 0.69 | 
Dilution gas flow (L/min) | 0.28 | 
Lens tune | Autotune | 
He flow (mL/min) | 1.0 | 5.0 (*10)
Energy discrimination (V) | 5.0 (*7.0) | 
Number of elements | 1 analyte, 1 ISTD | 27 analyte, 3 ISTD

## Slide 7: ISIS Parameters
ISIS 3 parameter | Setting | 
Loop Volume (mL) | 1.4 | 
 | Time (sec) | Pump Speed (%)
Sample load | 9 | 35
Stabilize | 23 | 5
Probe Rinse | 25 | 8
Probe Rinse 2 | 10 | 60
Probe Rinse 3 | 10 | 5
Optional Loop Probe Wash | 13 | 60
Optional Loop Wash | 3 | 5

## Slide 8: Experimental
Standards
  Agilent Environmental Calibration Standard (# 5183-4688) multi-element stock solution
  Single element standards for Al, Mn, Zn, Hg (Kanto Chemical), and P (SPEX)
  QC sample: Environmental Calibration Standard A (VHG)
  ISTD: Agilent Internal Standard Mix for ICP-MS (# 5188-6525)
Samples
  River Sediment A/ River Sediment B/ Estuarian Sediment/ Soil A/ Soil B
All standards and samples were prepared in 1% HNO3 and 0.5% HCl (HCl is beneficial for chemical stability of several critical trace elements, Hg, Ag, Mo, Tl…)
As QC samples, CCV and CCB were measured every after 10 samples.

## Slide 9: LOQs (10 sigma method)
*P and Se were measured by Enhanced He mode. As and Se were measured at the increased resolution settings used for M2+ Correction.** Pb was measured as the sum of the three most abundant isotopes, 206, 207 and 208.
 | He flow rate (mL/min) | ISTD | LOD (µg/kg)
9 Be | 1 | 6Li | 0.004
23 Na | 5 | 89Y | 7.75
24 Mg | 5 | 89Y | 1.02
27 Al | 5 | 89Y | 4.97
31 P* | 10 | 89Y | 2.75
39 K | 5 | 89Y | 30.3
44 Ca | 5 | 89Y | 17.1
51 V | 5 | 89Y | 0.016
52 Cr | 5 | 89Y | 0.371
55 Mn | 5 | 89Y | 0.222
56 Fe | 5 | 89Y | 3.27
59 Co | 5 | 89Y | 0.010
60 Ni | 5 | 89Y | 0.044
63 Cu | 5 | 89Y | 0.042
 | He flow rate (mL/min) | ISTD | LOD (µg/kg)
66 Zn | 5 | 89Y | 0.210
75 As* | 5 | 89Y | 0.049
78 Se* | 10 | 89Y | 0.049
95 Mo | 5 | 89Y | 0.009
107 Ag | 5 | 115In | 0.016
111 Cd | 5 | 115In | 0.008
121 Sb | 5 | 115In | 0.011
137 Ba | 5 | 115In | 0.059
201 Hg | 5 | 209Bi | 0.004
205 Tl | 5 | 209Bi | 0.005
Pb** | 5 | 209Bi | 0.163
232 Th | 5 | 209Bi | 0.008
238 U | 5 | 209Bi | 0.006
Notes: Note: LOQs are calculated using 10 sigma of 10 measurements of the blank (not 3 sigma).

## Slide 10: Long-term Stability Confirmed by ISTD Plot and QC Recovery
Chart: CCV recovery (%)
| ICV | CCV | CCV | CCV | CCV | CCV | CCV | CCV | CCV | CCV | CCV
9 Be | 102.1154 | 100.795 | 97.5742 | 100.7028 | 100.4004 | 100.1641 | 99.6653 | 98.176 | 98.7214 | 98.9506 | 97.9018
23 Na | 103.856 | 101.4383 | 97.1763 | 97.0472 | 97.1237 | 99.2933 | 101.5272 | 103.2343 | 103.2382 | 104.7486 | 105.4709
24 Mg | 104.3829 | 105.9575 | 101.1976 | 100.3034 | 100.4214 | 103.8598 | 103.9604 | 106.8681 | 105.46 | 107.4909 | 108.4786
27 Al | 101.393 | 106.2616 | 101.3694 | 99.8273 | 99.7161 | 101.5991 | 103.7218 | 103.9697 | 103.5368 | 106.6101 | 107.9389
31 P | 100.1902 | 104.5323 | 97.9169 | 99.6672 | 100.4213 | 98.3596 | 101.4386 | 105.5981 | 102.3064 | 98.8523 | 100.9603
39 K | 102.3459 | 102.738 | 103.5406 | 104.8493 | 104.6876 | 101.0398 | 100.1258 | 108.3939 | 107.0868 | 106.8243 | 108.5581
44 Ca | 108.5476 | 105.6758 | 102.9163 | 103.7196 | 103.3016 | 106.3969 | 106.9258 | 104.0796 | 103.9438 | 107.4718 | 106.7596
51 V | 99.7064 | 101.6794 | 102.6956 | 103.6277 | 103.8943 | 100.6867 | 99.1711 | 105.9344 | 104.3349 | 104.0867 | 103.8997
52 Cr | 103.7924 | 101.3024 | 101.7063 | 102.1989 | 101.6985 | 99.9404 | 97.8132 | 105.8272 | 102.2065 | 102.5026 | 101.7917
55 Mn | 100.6423 | 100.822 | 101.3847 | 102.4535 | 101.5389 | 98.6695 | 96.6149 | 103.9556 | 101.9246 | 101.903 | 101.9156
56 Fe | 105.9994 | 105.0191 | 105.0164 | 106.0508 | 105.4954 | 102.9293 | 100.3914 | 108.1266 | 105.9369 | 105.08 | 105.5033
59 Co | 101.7242 | 100.9657 | 96.5244 | 95.9091 | 95.6969 | 99.0646 | 98.2732 | 98.1403 | 97.7112 | 97.9539 | 96.2983
ISTD
QC sample recovery
8 hours
8 hours
Analysis of a mixed batch of CRMs, with QC block interspersed after every 10 samples
Notes: QC sample: every 10 samples
QC sample concentration: 50ppb for trace elements, 5000ppb for mineral elements, 500ppb for P, 1ppb for Hg

## Slide 11: Accurate Quantitative Results for All Elements in River Sediment A
(n=20) | Mean Conc | Certified Value | Recovery (%)
9 Be | 0.151 | - | -
23 Na | 51.8 | 50.0±0.3 | 103
24 Mg | 69.6 | 70.0±0.4 | 99
27 Al | 244 | 250±1 | 98
31 P | 0.017 | - | -
39 K | 149 | 150±1 | 99
44 Ca | 292 | 300±2 | 98
51 V | 0.252 | 0.250±0.003 | 101
52 Cr | 298 | 300±2 | 99
55 Mn | 7.74 | 8.00±0.04 | 97
56 Fe | 1190 | 1200±6 | 99
59 Co | 100 | 100±1 | 100
60 Ni | 0.51 | 0.50±0.01 | 102
63 Cu | 1.01 | 1.00±0.01 | 101
(n=20) | Mean Conc | Certified Value | Recovery (%)
66 Zn | 13.7 | 15.0±0.2 | 91
75 As | 0.594 | 0.60±0.03 | 99
78 Se | 19.7 | 20±2 | 98
95 Mo | 0.592 | - | -
107 Ag | 1.63 | - | -
111 Cd | 102 | 100±2 | 102
121 Sb | 518 | 500±20 | 104
137 Ba | 0.514 | 0.50±0.01 | 103
201 Hg | 0.028 | - | -
205 Tl | 9.72 | 10±1 | 97
206,207,208 Pb | 7.12 | 7.00±0.07 | 102
232 Th | 19.9 | 20±2 | 99
238 U | 9.82 | 10±1 | 98
All certified elements measured in the five sediment and soil CRMs were within ±10%
Concentration units mg/L except for elements in light blue rows which are in μg/L.

## Slide 12: Accurate Quantitative Results for All Elements in River Sediment B
(n=20) | Mean Conc | Certified Value | Recovery (%)
9 Be | 0.103 | - | -
23 Na | 54.2 | 50.0±0.5 | 108
24 Mg | 127 | 120±1 | 106
27 Al | 594 | 600±4 | 99
31 P | 9.72 | 10.0±0.1 | 97
39 K | 199 | 200±3 | 100
44 Ca | 294 | 300±2 | 98
51 V | 1.03 | 1.00±0.02 | 103
52 Cr | 15.8 | 15.0±0.1 | 105
55 Mn | 5.92 | 6.0±0.1 | 99
56 Fe | 410 | 400±2 | 102
59 Co | 162 | 150±3 | 108
60 Ni | 0.55 | 0.50±0.01 | 110
63 Cu | 1.06 | 1.00±0.01 | 106
(n=20) | Mean Conc | Certified Value | Recovery (%)
66 Zn | 4.7 | 5.00±0.05 | 94
75 As | 0.2 | 0.20±0.01 | 100
78 Se | 9.44 | 10±1 | 94
95 Mo | 2.3 | - | -
107 Ag | 0.318 | - | -
111 Cd | 30.6 | 30±1 | 102
121 Sb | 41 | 40±2 | 102
137 Ba | 4.02 | 4.00±0.04 | 100
201 Hg | 0.021 | - | -
205 Tl | 9.76 | 10±1 | 98
206,207,208 Pb | 2.04 | 2.0±0.1 | 102
232 Th | 97.8 | 100±5 | 98
238 U | 29 | 30±2 | 97
All certified elements measured in the five sediment and soil CRMs were within ±10%
Concentration units mg/L except for elements in light blue rows which are in μg/L.

## Slide 13: Accurate Quantitative Results for All Elements in Estuarian Sediment
(n=20) | Mean Conc | Certified Value | Recovery (%)
9 Be | 18.7 | 20.0±0.6 | 93
23 Na | 206 | 200±2 | 103
24 Mg | 102 | 100.0±0.6 | 102
27 Al | 710 | 700±4 | 102
31 P | 4.86 | 5.00±0.05 | 97
39 K | 146 | 150±2 | 97
44 Ca | 79 | 80.0±0.5 | 99
51 V | 1 | 1.00±0.02 | 100
52 Cr | 0.832 | 0.800±0.008 | 104
55 Mn | 3.86 | 4.00±0.04 | 96
56 Fe | 348 | 350±2 | 99
59 Co | 108 | 100±2 | 108
60 Ni | 0.32 | 0.30±0.02 | 107
63 Cu | 0.214 | 0.200±0.003 | 107
(n=20) | Mean Conc | Certified Value | Recovery (%)
66 Zn | 1.37 | 1.50±0.03 | 92
75 As | 0.097 | 0.100±0.005 | 97
78 Se | 45.8 | 50±5 | 92
95 Mo | 1.06 | - | -
107 Ag | 0.182 | - | -
111 Cd | 0.446† | 0.4 | -
121 Sb | 4.42† | 4 | -
137 Ba | 0.004 | - | -
201 Hg | 0.021 | - | -
205 Tl | 0.066 | - | -
206,207,208 Pb | 0.304 | 0.30±0.02 | 102
232 Th | 97.6 | 100±5 | 98
238 U | 0.055 | - | -
All certified elements measured in the five sediment and soil CRMs were within ±10%
†:Reference value
Concentration units mg/L except for elements in light blue rows which are in μg/L.

## Slide 14: Accurate Quantitative Results for All Elements in Soil A
(n=20) | Mean Conc | Certified Value | Recovery (%)
9 Be | 0.063 | - | -
23 Na | 72.2 | 70.0±0.7 | 103
24 Mg | 70.2 | 70.0±0.4 | 100
27 Al | 492 | 500±3 | 98
31 P | 9.84 | 10.0±0.1 | 98
39 K | 197 | 200±3 | 98
44 Ca | 344 | 350±2 | 98
51 V | 0.1 | 0.100±0.005 | 100
52 Cr | 0.016 | - | -
55 Mn | 0.104 | 0.100±0.001 | 104
56 Fe | 199 | 200.0±1.2 | 99
59 Co | 2.46 | - | -
60 Ni | 0.298 | 0.30±0.02 | 99
63 Cu | 0.3 | 0.300±0.005 | 100
(n=20) | Mean Conc | Certified Value | Recovery (%)
66 Zn | 0.916 | 1.00±0.002 | 92
75 As | 0.19 | 0.20±0.01 | 95
78 Se | 9.42 | 10±1 | 94
95 Mo | 0.234 | - | -
107 Ag | 0.236 | - | -
111 Cd | 2.98† | 6 | -
121 Sb | 32.6 | 30±1 | 109
137 Ba | 5.06 | 5.00±0.05 | 101
201 Hg | 0.019 | - | -
205 Tl | 0.066 | - | -
206,207,208 Pb | 0.4 | 0.40±0.02 | 100
232 Th | 97.6 | 100±5 | 98
238 U | 9.54 | 10±1 | 95
All certified elements measured in the five sediment and soil CRMs were within ±10%
Concentration units mg/L except for elements in light blue rows which are in μg/L.

## Slide 15: Accurate Quantitative Results for All Elements in Soil B
(n=20) | Mean Conc | Certified Value | Recovery (%)
9 Be | 0.398 | - | -
23 Na | 99.8 | 100±1 | 100
24 Mg | 79 | 80.0±0.48 | 99
27 Al | 716 | 700±4 | 102
31 P | 9.66 | 10.0±0.1 | 97
39 K | 206 | 210±3 | 98
44 Ca | 122 | 125±1 | 98
51 V | 0.788 | 0.800±0.016 | 99
52 Cr | 0.404 | 0.400±0.005 | 101
55 Mn | 97.4 | 100±1 | 97
56 Fe | 346 | 350±2 | 99
59 Co | 98.6 | 100±2 | 99
60 Ni | 0.202 | 0.20±0.01 | 101
63 Cu | 2.96 | 3.00±0.03 | 98
(n=20) | Mean Conc | Certified Value | Recovery (%)
66 Zn | 64 | 70.0±0.7 | 91
75 As | 5.74 | 6.00±0.12 | 96
78 Se | 0.254 | - | -
95 Mo | 1.55 | - | -
107 Ag | 0.5 | - | -
111 Cd | 195 | 200±4 | 97
121 Sb | 398 | 400±12 | 100
137 Ba | 6.98 | 7.00±0.07 | 100
201 Hg | 0.02 | - | -
205 Tl | 0.13 | - | -
206,207,208 Pb | 58.6 | 60.0±0.6 | 98
232 Th | 97.4 | 100±5 | 97
238 U | 242 | 250±10 | 97
All certified elements measured in the five sediment and soil CRMs were within ±10%
Concentration units mg/L except for elements in light blue rows which are in μg/L.

## Slide 16: Matrix Spike Recovery Test of Soil A
 | Soil-A (µg/kg) | Soil-A+ Spike 1 (µg/kg) | Soil-A+ Spike 2 (µg/kg) | Spike 1 Recovery (%) | Spike 2 Recovery (%) | Relative % Difference
9 Be | 0.18 | 92.1 | 90.8 | 92 | 91 | 1.1
51 V | 47.4 | 156 | 158 | 108 | 111 | 2.7
52 Cr | 1.38 | 104 | 107 | 102 | 105 | 2.9
55 Mn | 52 | 156 | 163 | 104 | 111 | 6.5
59 Co | 1.09 | 96.8 | 94.2 | 96 | 93 | 3.2
60 Ni | 157 | 264 | 268 | 107 | 112 | 4.6
63 Cu | 157 | 266 | 266 | 109 | 109 | 0
66 Zn | 524 | 691 | 688 | 111 | 109 | 1.8
75 As | 93.1 | 184 | 184 | 91 | 91 | 0
78 Se | 4.41 | 99.7 | 98.6 | 95 | 94 | 1.1
Matrix spikes were mostly within ±10%.

## Slide 17: Matrix Spike Recovery Test in Soil A ... continued
Matrix spikes were within ±10%.
 | Soil-A (µg/kg) | Soil-A+ Spike 1 (µg/kg) | Soil-A+ Spike 2 (µg/kg) | Spike 1 Recovery (%) | Spike 2 Recovery (%) | Relative % Difference
95 Mo | 0.1 | 100 | 97.8 | 100 | 98 | 2
107 Ag | 0.25 | 97.2 | 91.3 | 97 | 91 | 6.4
111 Cd | 1.43 | 103 | 102 | 101 | 101 | 0
121 Sb | 16.5 | 119 | 120 | 103 | 104 | 1
201 Hg | 0.01 | 2.06 | 2.08 | 102 | 103 | 1
205 Tl | 0.03 | 103 | 103 | 103 | 103 | 0
Pb* | 203 | 304 | 302 | 101 | 99 | 2
232 Th | 48.9 | 152 | 153 | 104 | 104 | 0
238 U | 4.82 | 107 | 108 | 102 | 103 | 1
*Pb was measured as the sum of the three most abundant isotopes, 206, 207 and 208.

## Slide 18: Conclusion
The Agilent 7850 ICP-MS accurately measured 28 elements in varied soil and sludge samples – typical, but difficult environmental samples.
A robust plasma (low CeO+/Ce+ ratio) ensured that variable soil sample matrices can be analyzed reproducibly over 8 hours, while ensuring consistency by minimizing matrix suppression.
The 7850 avoided common time-consuming and unproductive activities in the analytical workflow by predefining and automating many functions using a method wizard and autotune.
Polyatomic and M2+ interferences were removed using helium/KED and half-mass correction.
ISIS 3 provided high sample throughput and minimized deposits on the interface.
EMF and a post-run check reduced unnecessary maintenance, further improving productivity.

## Slide 19: More Information
Agilent Application Note
Agilent application note: 5994-2933EN